## Common Problems in FAAS and Their Solutions

| No. | Problem | Possible Causes | Solutions / Remedies |
| --- | --- | --- | --- |
| 1 | No absorbance signal (flat baseline) | • Lamp off, burned out, or wrong element lamp• Wrong wavelength setting• Nebulizer or burner blocked• Flame off or unstable | • Check lamp, alignment, and current• Set correct wavelength for element• Clean nebulizer/burner• Reignite and stabilize flame |
| 2 | Very low absorbance (weak signal) | • Sample concentration too low• Incorrect burner height or fuel/oxidant ratio• Instrument not optimized | • Increase sample concentration• Adjust flame settings• Optimize burner height and slit width |
| 3 | High background absorbance | • Molecular absorption or scattering• Organic matter or particulate contamination• Wrong flame type | • Use background correction (D₂ lamp)• Filter or digest samples• Use appropriate oxidizing/reducing flame |
| 4 | Unstable or noisy signal | • Flame flickering or gas pressure fluctuation• Air bubbles in aspiration line• Dirty nebulizer or burner• Poor lamp stability | • Stabilize gas flows• Remove air bubbles• Clean nebulizer/burner• Warm up lamp for 15–30 min |
| 5 | Nonlinear calibration curve | • Standards not prepared accurately• Concentration range too wide• Matrix interference | • Prepare fresh accurate standards• Work within linear range• Use matrix matching or standard addition |
| 6 | Signal drift during analysis | • Lamp temperature changes• Burner or nebulizer clogging• Changing gas flows | • Allow lamp to warm up• Clean burner and nebulizer• Maintain constant gas pressure |
| 7 | Carryover between samples | • Residual analyte in tubing or nebulizer• Inadequate rinsing | • Flush system with 1–2% HNO₃• Use blank between high and low samples |
| 8 | Interference from other elements (matrix effects) | • Chemical or spectral interference• High salt content or organics in sample | • Use background correction• Apply standard addition method• Dilute sample or use releasing agents |
| 9 | High blank reading | • Contaminated glassware, reagents, or water• Airborne dust contamination | • Use acid-washed glassware• Use ultrapure reagents and blanks• Work in clean area |
| 10 | Poor precision (high RSD) | • Inconsistent aspiration• Incomplete digestion• Inhomogeneous sample | • Check pump tubing• Digest sample completely• Homogenize sample before analysis |
| 11 | Flame goes out suddenly | • Wrong gas ratio• Blocked burner slot• Sudden pressure change | • Adjust gas flow (air–acetylene ratio)• Clean burner• Ensure steady pressure |
| 12 | Lamp current instability | • Old lamp or loose connection | • Replace hollow cathode lamp• Check electrical connections |
| 13 | Optical alignment problem | • Misaligned lamp, mirrors, or monochromator | • Realign optics according to manual• Perform wavelength calibration |
| 14 | Detection limit too high | • High background noise• Poor sensitivity• Excessive dilution | • Optimize conditions• Clean optics• Reduce dilution or use GFAAS |
| 15 | Incorrect results / bias | • Calculation errors• Uncorrected blank• Contaminated standards | • Double-check calculations• Always subtract reagent blank• Prepare fresh standards |

Tips for Reliable FAAS Operation

- Always warm up lamps before measurement (15–30 minutes).

- Optimize flame for each element — oxidizing or reducing flame.

- Run blanks and CRMs (Certified Reference Materials) regularly.

- Clean nebulizer and burner after each day’s use.

- Store standards in acidified bottles (1% HNO₃) to prevent precipitation.